## Sheet1
Technical Specififications of PPE-s and Medical Respiratory Items
LOT 1
No. | Item | Technical Specifications | Unit | Total Quantity
1 | Maska mjekësore/ kirurgjikale rezistente ndaj lengjeve /Fluid-resistant medical/surgical mask | High fluid resistance  Good breathability  Internal and external faces should be clearly identified  Structured design that does not collapse against the mouth (e.g. duckbill, cup shape)  Quality compliant with the below standards, or equivalent: - EN 14683 Type IIR performance - ASTM F2100 level 2 or level 3 or equivalent | piece | 67'500
2 | Doreza nitrili/Nitrile Gloves | Disposable Nitrile Gloves. Non sterile. Latex-free, odour-free and powder-free. Outer gloves preferably reach mid-forearm (minimum280mm total length). Different sizes. Quality compliant with the below standards, or equivalent:-EU standard directive 93/42/EEC Class I, EN 455-EU standard directive 89/686/EEC Category III, EN 374-ANSI/ISEA 105-2011-ASTM D6319-10. Box = 100 pieces | box | 34
3 | Termometer me ekran/Thermo scanner | Non Contact Infrared Forehead Thermometer, Digital with Fever Alarm and Memory Function, Instant Accurate Reads in 1 second°F and °C with Digital Display for Baby and Adults | piece | 90
4 | Barriere mbrojtese plastike per fytyren/Face shield mask | Safety Face Shield made of transparent recyclable PET, lightweight,fast and easy to wear.A large area from the eyebrows to the chin to protect against foreign aggression on your face. Has a foam strip, and the elastic band can be easily adjusted to suit habits, and it is firmly fixed and comfortable to wear. | piece | 340
5 | Veshje e padepertueshme, shumeperdorimshe qe pastrohet/ Isolation Gowns Washable | Isolation Gowns Washable Reusable. Material 99% Poly 1% Carbon; Complete fluid Protection- Preferally Color blue; Medium and Large size | piece | 680
6 | Pulsoksimeter/Finger Pulse Oximeter | Finger Pulse Oximeter, Blood Oxygen Saturation Monitor with OLED Screen Display with SPO2 and heart rate. | piece | 237
LOT 2
1 | Aparat oksigjenimi elektrik/Portable oxygen concentrator | Electricity powered device that concentrates oxygen from ambient air and delivers the concentrated oxygen in a controlled manner to a patient requiring oxygen therapy, 5-10 L/min. Device concentrates oxygen from ambient air On 4 antistatic swivel castors, 2 with brakes Integrated handle allows for easy moving and positioning Device weight: < 27 kg Provides a continuous flow of concentrated oxygen (> 82%) (preferably > 90%) from room air through one oxygen outlet. Oxygen sensing device is integrated and measures concentration at flow meter entrance. Oxygen outlet durable, with approx.6mm barbed fitting or equivalent Flowmeter range: 1 to 10 LPM Flowmeter continuously adjustable with markings at 0.5 L intervals Maximum output flow: 10 LPM Output pressure: min 55 // kPa. 0.04/0.07 Continuous monitoring, with visual and audible alert on: Low/ high output pressure, no flow, low oxygen concentration <82%, power failure Sound level produced: less than 52 dB (A) Power requirements: 100–240 V/50–60 Hz Mains power cable length ≥ 2.5 m. User interface easy to operate with numbers and displays that are clearly visible Digital or analogue meter to display cumulative hours of device operation Compliance to the following international standards (IEC 60601-1-11:2010 or equivalent). Warranty not less than 3 years | piece | 54