## Sheet1
Lista e pajisjeve per kujdesin ne banese/ List of equipment for home care LOT 1
Nr | Artikulli/Item | Specifikimet teknike/ Technical Specifications Requested | Sasia / Quantity | Cmimi /njesi/pa TVSH Price /Unit without VAT | Vlera pa tvsh/Value without VAT
1 | Aparat oksigjenimi elektrik/Portable oxygen concentrator | Electricity powered device that concentrates oxygen from ambient air and delivers the concentrated oxygen in a controlled manner to a patient requiring oxygen therapy. Maximum output flow: 5-10 LPM On 4 antistatic swivel castors, 2 with brakes; Integrated handle to allow easy moving and positioning Provides a continuous flow of concentrated oxygen (> 82%) (preferably > 90%) from room air through one oxygen outlet. Oxygen sensing device is integrated and measures concentration at flow meter entrance Flowmeter continuously adjustable with markings at 0.5 L intervals Output pressure: min 55 // kPa Continuous monitoring, with visual and audible alert on: Low/ high output pressure, no flow, low oxygen concentration <82%, power failure sound level produced: less than 52 dB (A) Power requirements: 220-230 V/50-60 Hz . Power cable length ≥ 2.5 m User interface easy to operate with numbers and displays that are clearly visible. Digital or analogue meter to display cumulative hours of device operation. Compliance to the following international standards; Device weight: Suitable for use in the patient houses. Warranty not less than 3 years. | 27 |  | 0
2 | Aparat EKG portative/ Portable ECG Machine | 12 Channel ECG Machine Electrocardiograph, 25cm color touch LCD Available for roll and Z-fold paper. Paper included. Accurate pacemaker pulse recognition and automatic analysis function. Working modes: manual, auto, storage. ECG waveform should be playbacked and frozen before or after 12 seconds. AC/DC power applicable, built in large capacity rechargeable; at least 250 cases storage . Voltage range:220/230V,50/60HZ; USB port interface for net transmission); Device weight: Suitable for use in the patient houses. Compliance to the following international standards; Warranty not less than 3 years . | 27 |  | 0
Total pa TVSH ne Leke/ Total without VAT, in ALL |  |  |  |  | 0
TVSH/ VAT |  |  |  |  | 0
Total meTVSH ne Leke/ Total with VAT in ALL |  |  |  |  | 0